# Analysis and validation of the different Brewer-Umkehr codes

## Redondas, E. Cuevas, V. Carreño and J. Sancho

Izaña Atmospheric Observatory (Instituto Nacional de Meteorología, Spain)

Contribution to the final Report of REVUE ("REconstruction of Vertical ozone distribution from Umkehr Estimates") Project, EC contract:    ENV4-CT95-0161.

February 2000

1.    Introduction

Preliminary comparisons of Umkehr profiles obtained with the Brewer spectrophotometer with other ozone instruments show good results (McElroy and Kerr, 1995). However the Umkehr code for the Brewer needs further investigation and validations because it is not so well tested that the algorithm used for Dobson instruments.

The Izaña Atmospheric Observatory is located at 28º18'N, 16º29'W on the island of Tenerife (Canary Islands) at 2367 m.a.s.l. The Station is managed by the Instituto Nacional de Meteorología (INM), Spain. The station is situated on the top of a mountain plateau on the dorsal ridge that crosses the island. The Observatory is normally above a temperature inversion layer, generally well established over the island (between 1200 and 1800 m a.s.l.), so air is free from local anthropogenic source influences and very often the Observatory is also free of clouds, resulting to be extremely clean and suitable for radiation measurements.

Continuous monitoring of total ozone started in May 1991 with a Brewer spectrophotometer (#33). The program of vertical ozone profiles was initiated in November 1992, with ozonesondes launches from Santa Cruz de Tenerife station (36 m.a.s.l.), distant 28 Km. from Izaña Observatory, on a weekly basis.

The main objective of the Izaña Atmospheric Observatory (IAO) group in the REVUE project is basically to validate the Umkehr algorithm used by the Brewer spectrophotometer against the ECC profiles under different sky conditions.

2. Experimental

Confident Umkehr profiles are obtained with the Brewer#33 at Izaña since September 1992. In August 1998 a new double Brewer spectrophotometer (#157) was installed at Izaña which was running in parallel with the older Brewer #33 till November 1998, when the #33 was moved to a new Spanish station.

Information concerning technical features and observation program capabilities with a Brewer Spectrophotometer can be found in McElroy et al., 1989, and in WMO/GAW, 1990. The Umkehr records are scheduled, on a daily basis, in the computer in the AM and PM 96º-62º zenith angle range.

ECC-6A-type ozonesondes, Scientific Pump Co. and RS-80 radiosondes (Väisäla) are flown with helium-filled rubber balloons (1200 g). A constant mixing ratio above burst level is assumed for the determination of the residual ozone if an altitude equivalent to 17 hPa has been reached. The integrated ozone column, for those ozone sondes that reach a burst level of at least 17 hPa, is ratioed with the same column amount measured with Brewer (starting at 2400 m a.s.l.). Most of the Brewer/ECC coefficients fall in the 0.95-1.05 range. ECC vertical profiles were corrected by the factor Brewer-O3/ECC-O3 and integrated for the different Umkehr layers (DU) shown in table 1. As Izaña station is at 2367 m .a.s.l., layers 0 and 1 are considered as an unique layer 1, in which the ozone integrated from sea level to Izaña’s altitude is subtracted. This amount is about 10 DU (Redondas et al., 1997). The ECC-Umkehr validation can be only performed for the first seven layers.

| Umkehr layer number | Pressure-layer base (hPa) | Altitude-layer base (km) |
| --- | --- | --- |
| 0 | 1000 | 0 |
| 1 | 500 | 5.5 |
| 2 | 250 | 10.3 |
| 3 | 125 | 14.7 |
| 4 | 62.5 | 19.1 |
| 5 | 31.2 | 23.5 |
| 6 | 15.6 | 28.0 |
| 7 | 7.81 | 32.6 |
| 8 | 3.91 | 37.5 |
| 9 | 1.95 | 42.6 |
| 10 | 0.98 | 47.9 |
| 11 | 0.49 | 53.2 |
| 12 | 0.24 | 58.3 |
| 13 | 0.12 | 63.1 |
| 14 | 0.061 | 67.8 |
| 15 | 0.030 | 72.2 |

Table 1; Pressure (hPa) an altitude (km) at the base of the different Umkehr layers. Pressure and altitude range where it is possible to perform an ECC-Umkehr validation is shaded.

Intensive ozonesounding campaigns carried out in July 1996 and 1997, in the framework of REVUE, have been performed to validate the Brewer-Umkehr codes.

A 3-wavelength (368, 500 and 778 nm.) PMOD/WRC sun-photometer was used to obtain the aerosol optical depth (AOD) for case analysis. Visual observations such as visibility, cloudiness (type, height and amount) and significant meteorological phenomena (haze, dust, fog, etc.) were recorded every 3 hours at the synoptic Izaña station and used as preliminary quality control criteria.

3. Validation of different ozone first-guess and Brewer-Umkehr codes.

One of the main problems experienced in the past when evaluating ECC-Umkehr comparisons was the fact that different versions of the Brewer-Umkehr code were used and that we did not know in detail the differences among them. So, the first task in this final period of REVUE was to careful analyse and re-compile the three existing Brewer-Umkehr codes in 32 bits.

Three Brewer-Umkehr codes are now available for the Brewer community: Wave 5-(1991) code, Wave 6 and Wave 5 (1994) codes, all of them based on the algorithm developed by Mateer and Deluisi (Mateer and DeLuisi, 1992).

The main features of the codes and the main differences among them are the following:

- The Wave-5 (91) code does not include the factor PFAC and the surface ozone content. This code scales the first guess to be in total ozone-compatible values. However the total ozone is not really right because the ozone coefficient appropriate to ozone-weighted mean temperature is not used. The temperature is assumed to be –44º C.

- The observed ozone on the Wave-5 and Wave-6 1994 codes is not a correct input value since the ozone is scaled using the ozone first guess and the temperature profile to –44º C. So, the observed ozone value in the algorithm depends on the chosen first guess.

Once it was checked that the new 32-bit codes gave the same results that early code versions, different tests and validations were performed. These are hereafter shown.

Three versions of ozone first-guess were also available: the original first-guess from Matter and De Luisi    (1992), the new one supplied by Irina Petropavlovskikh (personal communication), and the third one obtained for the Izaña station from the ECC climatology of Tenerife. This last algorithm uses four climatological profiles corresponding to the ozone sondes performed in December-February, March-May, June-August and September-November periods, respectively from November 1992 to December 1997 in Tenerife. These profiles provide information of the vertical ozone distribution from sea level to 32 km. For higher altitudes a constant mixing ratio is assumed.

3.1. Test of IPNOT and PFAC parameters. Case analysis: July 10th, 1997.

In order to evaluate the parameters controlling the scattering in the Brewer-Umkehr code some tests have been performed. IPNOT parameter is chosen in the code according to the altitude of the station. The best results were obtained with the default program value (IPNOT=1).

The Umkehr code assumes that the multiple scattering is proportional to the station pressure. PFAC an exponential factor of the pressure, with a default value of 1. As Izaña is a high mountain station, Tom McElroy (personal communication) suggested us to increase this value since a greater PFAC would take into account the multiple scattering enhancement from lower levels, circumstance that is specially important at Izaña because this station is normally above a permanent and strong sea cloud.

Figures 1 a-f, show the results of the Wave-6 code with the new first-guess of 1999 for July 10th (day 191 of 1997), with PFAC values of 1, 2, 0 and 3, respectively. These figures show for each day the first-guess, AM and PM Umkehr profiles and the ECC vertical profile integrated for each Umkehr layer. The percentage error (Umkehr-ECC)/ECC is also shown for Umkehr layers 1-6.

From different case analysis it can be observed that using PFAC=2 instead of default value PFAC=1, the comparisons with ECC profiles do not change significantly, the differences between the observed and calculated ozone decrease, but the RMSRES are higher and the number of iterations lightly increase. Similar results are obtained when using Wave-5 (94) as it can be derived from Figures 1 e-f. As an example are also shown the results for PFAC=0 and PFAC=3, respectively. It can be seen that these values of PFAC must not be used. Wave-5 (91) does not introduce the PFAC parameter.

The conclusion in this section is that PFAC=2 does not improve the results, so a PFAC=1 will be used in the future in the Wave-5 (94) and Wave-6 codes.

# 3.2. Considerations about the zenith angle range

The Brewer-Umkehr algorithm does not allow us to choose the zenith angle range in the Umkehr data process. The kbmin, kbmax, kemin, kemax parameters are only used in the algorithm to determine whether the Umkehr fits the range fixed by these parameters or not, and if it is, to process it. Any way, if the Umkehr is finally processed, it is done at the whole zenith angle range used in the measurement.

To obtain an estimation of the influence that on results could have a change in the zenith angle range, a laborious work should be done acting directly on the udu*.* files and reprocessing the Umkehr code, which will be performed in a near future. Any way, this test is of secondary importance.

3.3. Ozone covariance matrix test.

Several tests about the utilisation of the Uniform or the Climatological ozone distribution convarianze matrix have been performed. In Figure 2 the results on day July 10t, 1997 (191-97) are shown. The conclusion is that only minor changes in the results were observed when the SX matrix is modified.

3.4. Test of Wave5-old, Wave5 and wave6 Brewer-Umkehr codes and different first-guess. Case analysis: Jun-11th-97and Jul-10th-97

A comparison of the different Brewer-Umkehr codes has been performed analysing Umkehr data processed with the different Brewer-Umkehr codes, corresponding to ten pristine days chosen in different seasons of the years 1996 and 1997. In this report, only case analysis corresponding to June, 11th, 1997 and July, 10th, 1997 are shown in Figure 3 (a-i, June 11th; j-r July 10th) as an example. Strictly, this section constitutes a first-guess test more than a Brewer-Umkehr code comparison, as we will conclude from this study.

As it is derived from the analysis of the figures 3 a-r, the results obtained with different codes are quite similar each other, and the main differences in the results arise from the election of the first-guess.

It is difficult to recommend a Brewer-Umkehr code from the obtained results, however we will use the Wave-5 (91) for the following reasons:

a) Wave-5 (91) code is the most similar to the Dobson code.

b) The best available information is that of Wave-5 (91) code (Matter’s notes), (Interin Umkehr Inversion Algorithm for the Brewer Ozone Spectrophotometer).

c) In Wave5 (94) and Wave-6 there are several inputs which are not clear enough, such as surface ozone, IPNOT and PFAC parameters. Information about its physical meaning and utilisation is vague and poor.

d) The new instrumental error introduced in 1994 codes (Wave-5 and Wave-6) affects negatively the results.

It is worth to show in more detail the importance of the instrumental error. In previous REVUE reports we had used the instrumental error provided by default configuration in the 1994 codes. When we use the error vector provided in code Wave-5 (91) the results improve significantly, as it can be observed in the case analysis May 16th, 1997 (136-97), shown in Figure 4.

Plots of the old first-guess (1991), the new one provided by Irina Petropavlovskikh (personal communication), and that derived from the ozone sonde climatology of Tenerife, all of them used in the previous analysis, are shown in Figure 5 (a-c) for the year 1996. Significantly best results are obtained with the ozone climatology first-guess for all codes.

#### 4. Re-evaluation of the first and second REVUE intensive ECC-Umkehr intercomparisons.

The first and the second REVUE intensive ozonesounding campaigns were held on July 1996 and July 1997, respectively. ECC-Umkehr comparisons during these periods were shown in the first and second REVUE reports, respectively. In this final period a re-evaluation of these two intensive comparisons have been performed using Wave-5 (1991) old code and the Izaña ozone climatology first-guess.

4.1. July-96 Umkehr-ECC validation.

An ozonesonde (ECC-A6) intensive campaign was performed in July 1996 (from July 1st to July 31st) obtaining 22 quasi-simultaneous ECC-Umkehr profiles (Table 2). The results are shown in Figure 6 (a-u).

These analysis show that in general there is a good agreement between the ECC vertical profiles and the Umkehr profiles, with (Umk-ECC)/ECC errors of ±10% in Umkehr layers 2, 3, 4, 5 and 6. Even, for a significant number of analysed days and at some Umkehr layers the agreement ECC-Umkehr is within ±5%.

The wrong Umkehr in this campaign were the following:

Day 184 (both AM and PM).

Day 204 (both AM and PM, but only in layers 3 and 5).

Day 208 (both AM and PM).

Day 212 (PM)

Day 213 (both AM and PM).

All these Umkehr failed because cloudiness.

Some of the Umkehr showed higher errors than ±10%, only at some layers, due to unknown reasons.

Umkehr layers 1 and 2, the first one in the troposphere, and the second one with a major tropospheric part, do not use to agree the ECC profile.

Saharan intrusions, which produce high AOD and low visibility, do not affect much the Umkehr profiles, perhaps because these hazy air mass intrusions are confined to a relatively narrow layer in the lower troposphere (between 1 and 4 km altitude). For example, the ozone vertical profiles of ozone for July 19th and 20th are quite similar each other, both the ECC and the Umkehr profile. The ECC-Umkehr agreement is better in July 20th, however the AOD this day is one order of magnitude higher than the previous day.

It is considered that successful Umkehrs are characterized by a number of iterations (Iter) £ 3. Concerning the RMSRES (root-mean-square fit between observed Umkehr and Umkehr calculated from the retrieved profile), good Umkehr's should show RMSRES £ 1. These two parameters have been used as quality control criteria and are shown for the different days of the intensive campaign in Figure 7.

Another important quality control index is the count ratio of 323.2 nm and 303.6 nm wavelengths. Successful Umkehr's always show a smooth and well-fitted curve when this ratio is represented against the zenith angle. In Figure 8 a “good day” (190) and a “bad day” (213) are shown as an example. However the difference between the ozone calculated by the Umkehr and the ozone observed is not a good quality criteria. This difference is almost always very small, even when the Umkehr is not good and the SMSRES and Iter are out of acceptable values.

Once the wrong Umkehrs have been removed applying the mentioned quality control criteria, a picture of the (Umk-ECC)/ECC errors for the Umkehr layers 1, 2, 3, 4, 5 and 6 are shown for the whole intensive campaign in Figure 9.

As it can be observed in Figure 9 and in Table 3, layer 1 is underestimated by the Umkehr (18%). In the lower troposphere the Umkehr algorithm, as it is well known, does not provide good results (layer 1). In layer 2 the results are much better than initially expected, probably because the incorporation of the new ECC climatology based first-guess. In this layer is located the double tropopause typical of the subtropical region (Sancho et al., 1999), at 12 km and 16 km, respectively. The ozone climatology first-guess for Tenerife has improved significantly the results in this layer. Layers 3, 4 and 5 agree quite well the ECC. Most of the Umkehr-ECC differences fall in the range ± 10% (see table 3), being a significant number of them in the ± 5% interval, and only at layer 5 are found a few percentage (8.4%) of PM Umkehrs with errors exceeding 20%. The best results are obtained in layer 4. So, from 15 km to 28 km, region where is found the ozone peak in the subtropics (around 25 km altitude), the Umkehr algorithm provides reasonable good results. About 60% of the Umkehr-ECC differences fall in the ± 10% range in layer 6. However, about 20% of the Umkehr profiles show errors higher than 20%.

In this campaign the number of wrong Umkehr's was significantly higher at PM (10, 45%) than at AM (4, 18%).

| Natural day (1996) | Cloudiness | AOD |
| --- | --- | --- |
| 183 (Jul-1) | 0 | 0.1<AOD<1 |
| 184 (2) | 1/8 High | 0.1<AOD<0.5 |
| 185 (3) | ---- | ---- |
| 186 (4) | ---- | ---- |
| 187 (5) | ---- | ---- |
| 188 (6) | ---- | ---- |
| 189 (7) | ---- | ---- |
| 190 (8) | 0 | AOD<0.03 |
| 191 (9) | 0 | AOD<0.05 |
| 192 (10) | 0 | AOD<0.05 |
| 193 (11) | 0 | AOD<0.05 |
| 194 (12) | 0 | 0.05<AOD<0.1 |
| 195 (13) | 0 | 0.1<AOD<0.5 |
| 196 (14) | 0 | 0.1<AOD<0.5 |
| 197 (15) | 0 | 0.08<AOD<0.12 |
| 198 (16) | 0 | 0.1<AOD<0.07 |
| 199 (17) | 0 | AOD<0.03 |
| 200 (18) | ---- | ---- |
| 201 (19) | 0 | 0.08<AOD<0.12 |
| 202 (20) | 0 | 0.2<AOD<0.3 |
| 203 (21) | ---- | ---- |
| 204 (22) | 1/8 Low | 0.2<AOD<0.34 |
| 205 (23) | 0 | 0.2<AOD<0.4 |
| 206 (24) | 0 | 0.2<AOD<0.4 |
| 207 (25) | 0 | 0.1<AOD<0.3 |
| 208 (26) | 4/8 Mid | 0.3<AOD<0.5 |
| 209 (27) | ---- | ---- |
| 210 (28) | ---- | ---- |
| 211 (29) | 0 | 0.03<AOD<0.06 |
| 212 (30) | 4/8 Mid (PM) | 0.1<AOD<0.3 |
| 213 (31) | 4/8 Mid | 0.3<AOD<0.5 |

Table 2; Days of the first Umkehr-ECC intensive intercomparison performed in July 1996. Cloudiness and Aerosol Optical Depth (at 500 nm) are indicated in the second and third column, respectively.

| July-96 | [0%,±5%] (%) ### AM/PM | ### (5%,10%] ### (-5%,-10%] ### (%) ### AM/PM | ### (10%,20%] ### (-10%,-20%] ### (%) ###### AM/PM | ###### <-20% ###### >20% ###### (%) ###### AM/PM | ###### Mean error ###### (%) ###### AM/PM |
| --- | --- | --- | --- | --- | --- |
| Layer 1 | 16.8 / 8.3 | 5.5 / 8.3 | 22.2 / 33.4 | 55.5 / 50.0 | -18% /-19% |
| Layer 2 | 16.7 / 41.6 | 27.8 / 25 | 33.3 / 16.7 | 22.2 / 16.7 | +13% / +4.6% |
| Layer 3 | 44.4/ 50.0 | 16.7 / 16.7 | 38.9 / 33.3 | 0 / 0 | -3.4% / -4.9% |
| Layer 4 | 77.8 / 75.0 | 11.1 / 0 | 11.1 / 25.0 | 0 / 0 | +0.8% / -4.1% |
| Layer 5 | 57.9 / 25.0 | 26.3 / 33.3 | 15.8 / 33.3 | 0 / 8.4 | +4.8% / +8.3% |
| Layer 6 | 27.8 / 35.8 | 33.3 / 21.4 | 22.2 / 21.4 | 16.7 / 21.4 | +9.1%/ +15.4% |

Table 3; Percentage of Umkehr in different (Umk-ECC)/ECC*100 error ranges for Umkehr layers 1, 2, 3, 4, 5 and 6 for July 1996. Twenty-two (22) quasi-simultaneous ECC-Umkehr profiles were analysed. Wrong Umker's were discarded in this analysis.

4.2. July-97 Umkehr-ECC validation.

In this campaign a total of 25 ozone sondes (ECC-A6) were launched in July 1997 (from July 2nd to July 31st), as it is indicated in table 4. Quasi-simultaneous ECC-Umkehr profiles are shown in Figure 10 (a-z).

The Umkehr profiles in July 1997 were obtained with two Brewer’s:

- Brewer#70 (Madrid) in the period July 1st-16th. This instrument was running at Izaña during this period because the #33 was participating in the SUSPEN intercomparison, which was held at Thessaloniki at that time.

- Brewer#33 (Izaña) during the rest of the intercomparison.

This fact gave us the opportunity to check the Umkehr algorithm with two different instruments operating at the same site.

Only one Umkehr failed in this intensive campaign (AM Umkehr of Day 199). Umkehr's in days with significant amount of high and middle clouds succeeded, as well as Umker's performed during hazy days caused by Saharan intrusions.

The quality control criteria (number of iterations and the root-mean-square fit between observed Umkehr and Umkehr calculated from the retrieved profile (RSMRES) are shown in Figure 11. Only the Umkehr of day 199 AM shows a RSMRES too high (10) and 5 iterations. The Umkehr's corresponding to day 200 (AM), day 202 (AM) and day 204 (PM) show a number of iterations of 4, however these Umkehr profiles agree well the ECC profiles. Once the wrong Umkehr has been removed applying the mentioned quality control criteria, a picture of the (Umk-ECC)/ECC errors for the Umkehr layers 1, 2, 3, 4, 5 and 6 are shown for the whole intensive campaign in Figure 12.

| Natural day (1997) | Cloudiness | AOD |
| --- | --- | --- |
| 183 (Jul-2) | 0 | AOD=0.01 |
| 184 (3) | 0 | AOD=0.02 |
| 185 (4) | 0 | AOD=0.01 |
| 186 (5) | ---- | ---- |
| 187 (6) | ---- | ---- |
| 188 (7) | 0 | AOD=0.06 |
| 189 (8) | 0 | AOD=0.18 Dust observed |
| 190 (9) | 0 | AOD=0.18 Dust observed |
| 191 (10) | 0 | AOD=0.01 |
| 192 (11) | 6/8 High | AOD=0.01 |
| 193 (12) | 0 | AOD=0.08 |
| 194 (13) | ---- | ---- |
| 195 (14) | 0 | AOD=0.02 |
| 196 (15) | 0 | AOD=0.02 |
| 197 (16) | 0 | AOD=0.01 |
| 198 (17) | 4/8 High (PM) | AOD=0.24 Dust observed |
| 199 (18) | 0 | AOD=0.21 Dust observed |
| 200 (19) | 0 | AOD=0.02 |
| 201 (20) | ---- | ---- |
| 202 (21) | 0 | AOD=0.02 |
| 203 (22) | 0 | AOD=0.02 |
| 204 (23) | 0 | AOD=0.02 |
| 205 (24) | 0 | AOD=0.02 |
| 206 (25) | 0 | AOD=0.06 |
| 207 (26) | 4/8 Mid 4/8 High AM | AOD=0.16 Dust observed |
| 208 (27) | ---- | ---- |
| 209 (28) | 0 | AOD=0.3 Dust observed |
| 210 (29) | 0 | AOD=0.26 |
| 211 (30) | 0 | AOD=0.12 |
| 212 (31) | 0 | AOD=0.02 |

Table 4; Days of the second Umkehr-ECC intensive intercomparison performed in July 1997. Cloudiness and Aerosol Optical Depth (at 500 nm), as well as visual observations, are indicated in the second and third column, respectively.

As it can be observed in Figure 12 and in Table 5, layers 3, 4, 5 and 6 show more than 50% of data with associated errors within ± 10% range. During this campaign an especial good agreement was found in layers 4 and 5 in which more than 80% of data showed errors within ± 10%.    The tropospheric layer 1 showed unexpected good results during this campaign (around 50% of data showed errors within ± 10% range).

| July-97 | [0%,±5%] (%) ### AM/PM | ### (5%,10%] ### (-5%,-10%] ### (%) ### AM/PM | ### (10%,20%] ### (-10%,-20%] ### (%) ###### AM/PM | ###### <-20% ###### >20% ###### (%) ###### AM/PM | ###### Mean error ###### (%) ###### AM/PM |
| --- | --- | --- | --- | --- | --- |
| Layer 1 | 37.0/25.0 | 29.0/ 20.0 | 29.0 / 28.0 | 5.0 / 20.0 | -6.3% /-8.2% |
| Layer 2 | 4.2 / 8.2 | 12.5 / 4.2 | 29.2 / 33.3 | 54.1 / 58.3 | +27.5%/ +24.1% |
| Layer 3 | 45.8/ 36.0 | 4.2 / 20.0 | 33.3 / 32.0 | 16.7/ 12.0 | +27.5% / +24.1% |
| Layer 4 | 37.5 / 41.7 | 29.2 / 33.3 | 25.0 / 25.0 | 8.3 / 0 | +7.5% / +5.2% |
| Layer 5 | 65.3 / 29.2 | 26.0 / 62.5 | 8.7 / 8.3 | 0 / 0 | -2.1% / +0.6% |
| Layer 6 | 39.1 / 37.5 | 17.4 / 20.8 | 34.8 / 33.3 | 8.7 / 8.4 | -2.1%/ +1.3% |

Table 5; Percentage of Umkehr in different (Umk-ECC)/ECC*100 error ranges for Umkehr layers 1, 2, 3, 4, 5 and 6 for July 1997. Twenty-five (25) quasi-simultaneous ECC-Umkehr profiles were analysed. The wrong Umkehr of day 199 (AM) was discarded in this analysis.

The intensive ECC-Umkehr intercomparison performed in July-97 shows that the Umkehr profile detects well de ozone maximum centred on layers 4 and 5. Most of the 25 ECC-Umkehr intercomparison performed in this campaign showed good agreement.

#### 5. Main conclusions

- Different tests performed with the quasi-simultaneous ECC-Umkehr profiles corresponding to the two intensive campaigns of July 1996 and July 1997, show that the best results are obtained using the Wave5-old (1991) Umkehr code, and the ozone first-guess derived from the ozonesonde climatology of Tenerife station.

- The best values of the parameters controlling the scattering in the Brewer-Umkehr (IPNOT and PFACT) are those that Umkehr program uses as default.

- A good agreement between the first guess, AM and PM Umkehr's and ECC is observed in almost all the profiles compared during the first and the second REVUE intensive campaigns, mainly in layers 4, 5 and 6.

- The goodness of fit of a Umkehr profile is directly related, among other parameters, to the sky conditions. Intensive comparisons carried out in July-96 and July-97 show that, as expected, cloudiness affects very negatively to the Umkehr. However, a low amount of middle-high clouds and Saharan intrusions, with high aerosol optical depth, cause a negligible influence on the Umkehr profile retrieval.

- The number of iterations and the RMSRES are the most simple and helpful quality control parameters that can determine goodness of fit of a Umkehr profile.

Acknowledgments

The authors are especially grateful to Dr. Tom McElroy (AES, Canada) for providing crucial information of the Brewer Umkehr code.

The authors thanks Sergio Afonso who performed successfully the ozone sonde intensive campaigns of July 1996 and July 1997.

##### References

Mateer, C.L. and J.J.DeLuisi, A new inversion algorithm, J.Atm. and Terr. Phys., 54, 537-556, 1992.

McElroy, C.T., C.L. Mateer, J.B. Kerr, and D.I Wardle, 1989, Umkehr Observations made with the Brewer ozone Spectrophotometer, Ozone in the Atmosphere, A. Deepak Publishing, 1989.

McElroy, C.T. and J. Kerr, Table Mountain Ozone Intercomparison: Brewer Umkehr Observations, J. Geophys. Res., 1995.

Redondas, A., E. Cuevas, J. Sancho, Investigaciones sobre Ozono Estratosférico y Radiación Ultravioleta en el Observatorio de Vigilancia Atmosférica de Izaña, Física de la Tierra, nº 9, 29-66, Servicio de Publicaciones de la Universidad Complutense de Madrid, Editorial Complutense, 1997.

WMO/GAW, WMO Consultation on Brewer Ozone Spectrophotometer Operation, Calibration and Data Reporting, 1990.

##### Figures captions

Figure 1 (a-f); different results obtained for July 10th (day 191 of 1997) using different PFAC factors.

Figure 2; Results of the climatological ozone distribution convarianze matrix test for July 10th, 1997 (191-97).

Figure 3; Example of comparison of the different Brewer-Umkehr codes. The analysis corresponds to June, 11th, 1997 and July, 10th, 1997 (a-i, June 11th; j-r July 10th).

Figure 4; Results of the error vector test for the case analysis May 16th, 1997 (136-97).

Figure 5 (a-c); Plots of the old first-guess (1991), the new one provided by Irina Petropavlovskikh (personal communication), and that derived from the ozone sonde climatology of Tenerife, corresponding to 1996.

Figure 6 (a-u); Twenty two (22) quasi-simultaneous ECC-Umkehr profiles from the ozonesonde (ECC-A6) intensive campaign performed in July 1996 (from July 1st to July 31st).

Figure 7; RMSRES (root-mean-square fit between observed Umkehr and Umkehr calculated from the retrieved profile) and number of iterations (Iter) for the Umkehr's performed in July 1996 (AM and PM).

Figure 8; Count ratio of 323.2 nm and 303.6 nm wavelengths against the zenith angle for a “good day” (190) and a “bad day” (213).

Figure 9; (Umk-ECC)/ECC errors for the Umkehr layers 1, 2, 3, 4, 5 and 6 for the whole July 1996 intensive campaign.

Figure 10; Twenty five (25) quasi-simultaneous ECC-Umkehr profiles from the ozonesonde (ECC-A6) intensive campaign performed in July 1997 (from July 2nd to July 31st).

Figure 11; RMSRES (root-mean-square fit between observed Umkehr and Umkehr calculated from the retrieved profile) and number of iterations (Iter) for the Umkehr's performed in July 1997 (AM and PM).

Figure 12; (Umk-ECC)/ECC errors for the Umkehr layers 1, 2, 3, 4, 5 and 6 for the whole July 1997 intensive campaign.